Pressure correction table for flow measurement

on the float principle / gas measuring

pe=0…12bar

| New pressure Pe [bar] | Calibration Pressure pe [bar] |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | New pressure Pe [bar] |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 0 | 0,2 | 0,4 | 0,6 | 0,8 | 1 | 1,5 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 |  |
| 0 | 1 | 0,913 | 0,847 | 0,790 | 0,745 | 0,707 | 0,632 | 0,577 | 0,500 | 0,446 | 0,408 | 0,378 | 0,353 | 0,333 | 0,316 | 0,302 | 0,289 | 0,277 | 0 |
| 0,2 | 1,095 | 1 | 0,926 | 0,866 | 0,817 | 0,775 | 0,693 | 0,632 | 0,547 | 0,490 | 0,447 | 0,414 | 0,387 | 0,365 | 0,347 | 0,330 | 0,316 | 0,304 | 0,2 |
| 0,4 | 1,184 | 1,080 | 1 | 0,937 | 0,884 | 0,837 | 0,750 | 0,684 | 0,592 | 0,529 | 0,482 | 0,446 | 0,417 | 0,395 | 0,375 | 0,357 | 0,342 | 0,328 | 0,4 |
| 0,6 | 1,265 | 1,155 | 1,070 | 1 | 0,943 | 0,895 | 0,800 | 0,730 | 0,632 | 0,565 | 0,517 | 0,477 | 0,447 | 0,421 | 0,400 | 0,381 | 0,365 | 0,351 | 0,6 |
| 0,8 | 1,340 | 1,220 | 1,130 | 1,060 | 1 | 0,950 | 0,850 | 0,775 | 0,672 | 0,600 | 0,548 | 0,506 | 0,474 | 0,447 | 0,425 | 0,404 | 0,387 | 0,372 | 0,8 |
| 1,0 | 1,414 | 1,291 | 1,195 | 1,119 | 1,055 | 1 | 0,896 | 0,817 | 0,707 | 0,632 | 0,577 | 0,535 | 0,500 | 0,471 | 0,447 | 0,426 | 0,408 | 0,393 | 1,0 |
| 1,5 | 1,580 | 1,445 | 1,336 | 1,250 | 1,180 | 1,120 | 1 | 0,913 | 0,791 | 0,707 | 0,645 | 0,597 | 0,559 | 0,527 | 0,500 | 0,476 | 0,457 | 0,439 | 1,5 |
| 2 | 1,730 | 1,580 | 1,460 | 1,370 | 1,290 | 1,220 | 1,095 | 1 | 0,866 | 0,774 | 0,707 | 0,654 | 0,612 | 0,577 | 0,549 | 0,522 | 0,500 | 0,480 | 2 |
| 3 | 2,000 | 1,820 | 1,690 | 1,580 | 1,490 | 1,410 | 1,265 | 1,150 | 1 | 0,892 | 0,816 | 0,755 | 0,707 | 0,667 | 0,632 | 0,602 | 0,577 | 0,555 | 3 |
| 4 | 2,240 | 2,040 | 1,890 | 1,770 | 1,666 | 1,584 | 1,414 | 1,290 | 1,119 | 1 | 0,913 | 0,845 | 0,791 | 0,746 | 0,707 | 0,675 | 0,645 | 0,620 | 4 |
| 5 | 2,450 | 2,240 | 2,070 | 1,930 | 1,830 | 1,730 | 1,550 | 1,410 | 1,220 | 1,095 | 1 | 0,925 | 0,865 | 0,817 | 0,775 | 0,738 | 0,707 | 0,680 | 5 |
| 6 | 2,650 | 2,420 | 2,240 | 2,090 | 1,970 | 1,870 | 1,670 | 1,530 | 1,320 | 1,180 | 1,080 | 1 | 0,936 | 0,882 | 0,838 | 0,797 | 0,765 | 0,734 | 6 |
| 7 | 2,830 | 2,580 | 2,390 | 2,230 | 2,110 | 2,000 | 1,790 | 1,630 | 1,410 | 1,260 | 1,150 | 1,070 | 1 | 0,943 | 0,895 | 0,852 | 0,816 | 0,785 | 7 |
| 8 | 3,000 | 2,740 | 2,535 | 2,375 | 2,240 | 2,120 | 1,897 | 1,732 | 1,500 | 1,343 | 1,225 | 1,135 | 1,060 | 1 | 0,950 | 0,905 | 0,867 | 0,832 | 8 |
| 9 | 3,165 | 2,890 | 2,670 | 2,500 | 2,360 | 2,240 | 2,000 | 1,820 | 1,580 | 1,410 | 1,290 | 1,195 | 1,120 | 1,050 | 1 | 0,952 | 0,912 | 0,877 | 9 |
| 10 | 3,320 | 3,030 | 2,800 | 2,620 | 2,480 | 2,350 | 2,100 | 1,920 | 1,660 | 1,480 | 1,350 | 1,250 | 1,170 | 1,110 | 1,050 | 1 | 0,957 | 0,920 | 10 |
| 11 | 3,460 | 3,160 | 2,930 | 2,740 | 2,580 | 2,450 | 2,190 | 2,000 | 1,730 | 1,540 | 1,410 | 1,305 | 1,225 | 1,155 | 1,095 | 1,040 | 1 | 0,960 | 11 |
| 12 | 3,600 | 3,290 | 3,040 | 2,850 | 2,690 | 2,550 | 2,280 | 2,080 | 1,800 | 1,610 | 1,470 | 1,360 | 1,275 | 1,200 | 1,140 | 1,087 | 1,040 | 1 | 12 |

The above table is used to correct the flow meter for gases displayed values, if the operating pressure differs from that oft he underlying pressure calibration.

Example::

Calibration pressure  Pe = 5bar / Working pressure Pe = 2bar

Under Calibration pressure Pe = 5bar is read in line 8 under new pressure  Pe = 2bar from the factor 0,707 . The values displayed by the meter must be multiplied by.

The Factor „F“ will be:

F=new Pabscalibration Pabs = 3bar6bar=0,707

Attention:                                                      notice:

Pe = effective pressure                                   New pressure higher: more than 1 factor
                                                                        New pressure is smaller: a factor in
                                                                        Intermediate values ​​are interpolated

!